## Definitions
scopeTypes | dataCategoryType | measurandType | unit | equipmentClass | Statement categories | Equipment Categories | metaDataType
- | - | - | - |  | general | item | Data
reference | Value | Measure.Conductance | \metre |  | norm | reference | UsedReferenceId
itemIndication | TargetValue | Measure.Conductivity | \kilogram |  | method |  | UsedMethodId
itemBias | ToleranceLimitLower | Measure.Current.AC | \second |  | exception |  | UsedEquipmentId
environment | ToleranceLimitUpper | Measure.Current.AC.Sinewave | \ampere |  | attention |  | customerTag
| AcceptanceLimitLower | Measure.Current.AC.Squarewave | \kelvin |  | traceability |  | laboratoryTag
| AcceptanceLimitUpper | Measure.Current.AC.Trianglewave | \mole |  | comment |  | accreditationException
| Conformity | Measure.Current.DC | \candela |  | accreditation |  | TimeStamp
| repeatability | Measure.Density.Mass.Gas | \one |  | accreditationException |  | Exception
| Other | Measure.Density.Mass.Liquid | \day |  |  |  | Statement
| accreditationApplies | Measure.Density.Mass.Solid | \hour
| ExpandedUncertainty | Measure.Force | \minute
| UncertaintyCoverageFactor_k | Measure.Frequency | \degree
| UncertaintyCoverageProbability | Measure.Frequency.AmplitudeModulation.Rate | \arcminute
|  | Measure.Frequency.FrequencyModulation.Deviation | \arcsecond
|  | Measure.Frequency.FrequencyModulation.Rate | \gram
|  | Measure.Frequency.PhaseModulation.Rate | \radian
|  | Measure.Humidity.Absolute | \steradian
|  | Measure.Humidity.Relative | \hertz
|  | Measure.Impedance | \newton
|  | Measure.Inductance | \pascal
|  | Measure.Length | \joule
|  | Measure.Length.Circumference | \watt
|  | Measure.Length.Diameter | \coulomb
|  | Measure.Length.Form.Flatness | \volt
|  | Measure.Length.Form.Parallelism | \farad
|  | Measure.Length.Form.Perpendicularity | \ohm
|  | Measure.Length.Form.Roughness | \siemens
|  | Measure.Length.Form.Roundness | \weber
|  | Measure.Length.Form.Sphericity | \tesla
|  | Measure.Length.Form.Straightness.Axis | \henry
|  | Measure.Length.Form.Straightness.Surface | \degreecelsius
|  | Measure.Length.Radius | \lumen
|  | Measure.Mass.Apparent | \lux
|  | Measure.Mass.Conventional | \becquerel
|  | Measure.Mass.True | \sievert
|  | Measure.Phase.PhaseModulation | \gray
|  | Measure.Phase.ReflectionFactor.RF | \katal
|  | Measure.Phase.TransmissionFactor | \hectare
|  | Measure.PhaseNoise.SideBand | \litre
|  | Measure.Power.RF.Sinewave | \tonne
|  | Measure.Pressure.Hydraulic.Static | \electronvolt
|  | Measure.Pressure.Pneumatic.Absolute.Static | \dalton
|  | Measure.Pressure.Pneumatic.Differential.Static | \astronomicalunit
|  | Measure.Pressure.Pneumatic.Gage.Static | \neper
|  | Measure.Ratio.AmplitudeModulation | \bel
|  | Measure.Ratio.Density.Mass | \decibel
|  | Measure.Ratio.Humidity.Relative | \deca
|  | Measure.Ratio.Humidity.Specific | \hecto
|  | Measure.Ratio.Power.ReflectionFactor.RF | \kilo
|  | Measure.Ratio.Power.RF.Sinewave.Delta.Frequency | \mega
|  | Measure.Ratio.Power.RF.Sinewave.Delta.Power | \giga
|  | Measure.Ratio.Power.TransmissionFactor | \tera
|  | Measure.Ratio.Torque | \peta
|  | Measure.Ratio.Voltage.AC.Ripple.OnDC | \exa
|  | Measure.Ratio.Voltage.AC.Sinewave.Delta.Frequency | \zetta
|  | Measure.Ratio.Voltage.AC.Sinewave.Delta.Voltage | \yotta
|  | Measure.Resistance | \deci
|  | Measure.Temperature | \centi
|  | Measure.Temperature.Radiometric | \milli
|  | Measure.Temperature.Simulated.PRT | \micro
|  | Measure.Temperature.Simulated.RTD | \nano
|  | Measure.Temperature.Simulated.Thermocouple | \pico
|  | Measure.Time.Transition | \femto
|  | Measure.Time.UTC | \atto
|  | Measure.Torque | \zepto
Measure.Torque.HydraulicPressure
Measure.Voltage.AC
Measure.Voltage.AC.Ripple.OnDC
Measure.Voltage.AC.Sinewave
Measure.Voltage.AC.Squarewave
Measure.Voltage.AC.Trianglewave
Measure.Voltage.DC
Measure.Weight
Source.Capacitance
Source.Conductance
Source.Conductivity
Source.Current.AC.Sinewave
Source.Current.AC.Squarewave
Source.Current.AC.Trianglewave
Source.Current.DC
Source.Density.Mass.Gas
Source.Density.Mass.Liquid
Source.Density.Mass.Solid
Source.Force
Source.Frequency.AC.Squarewave
Source.Frequency.Arbitrary.Cardiograph
Source.Humidity.Absolute
Source.Impedance
Source.Inductance
Source.Length
Source.Length.Circumference
Source.Length.Diameter
Source.Length.Form.Perpendicularity
Source.Length.Form.Roundness
Source.Length.Form.Sphericity
Source.Length.Form.Straightness.Surrface
Source.Length.Radius
Source.Mass.Apparent
Source.Mass.Conventional
Source.Mass.True
Source.Period.Marker
Source.Period.Squarewave
Source.Power.AC.Sinewave
Source.Power.AC.Sinewave.Simulated
Source.Power.DC.Simulated
Source.Power.Noise.Terminated
Source.Power.RF.Sinewave
Source.Pressure.Hydraulic.Static
Source.Pressure.Pneumatic.Absolute.Static
Source.Pressure.Pneumatic.Differential.Static
Source.Pressure.Pneumatic.Gage.Static
Source.Ratio.AmplitudeModulation
Source.Ratio.Humidity.Relative
Source.Ratio.Humidity.Specific
Source.Ratio.Power.RF.Sinewave.Delta.Frequency
Source.Ratio.Voltage.AC.Sinewave.Delta.Frequency
Source.Ratio.Voltage.AC.Sinewave.Delta.Voltage
Source.Resistance
Source.Temperature
Source.Temperature.FixedPoint
Source.Temperature.Radiometric
Source.Temperature.Simulated.PRT
Source.Temperature.Simulated.RTD
Source.Temperature.Simulated.Thermocouple
Source.Time.Marker
Source.Torque
Source.Voltage.AC.Sinewave
Source.Voltage.DC
Source.Voltage.DC.Delta.Voltage
Source.Voltage.Noise.Terminated
Source.Voltage.Shorted
Measure.Volume
miscelaneous01
miscelaneous02
miscelaneous03
miscelaneous04
miscelaneous05

## AdministrativeData
heading lang1 | heading lang2 | Description | Value | XPATH
Calibration certificate | Kalibreringscertifikat | Document_name |  | /dcc:digitalCalibrationCertificate
This document was created with: | Dette dokument blev genereret ved hjælp af: |  |  | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:dccSoftware/dcc:software
|  | Software_Name | excel2dcc.py | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:dccSoftware/dcc:software/dcc:name
|  | Software_Version | v0.0.1 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:dccSoftware/dcc:software/dcc:release
|  |  | DK | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:countryCodeISO3166_1
|  |  | da | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:usedLangCodeISO639_1
|  |  | en | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:mandatoryLangCodeISO639_1
Certificate number | Certifikatnummer |  |  | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:uniqueIdentifier
|  | Unique_ID | SKH_10112_2 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:uniqueIdentifier/dcc:value
Order number | Ordrenummer |  |  | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:customerIdentification
|  | OrderNo | xx28 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:customerIdentification/dcc:value
|  | ArrivalDate | 2022-02-09 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:receiptDate
|  | StartDate | 2022-02-10 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:beginPerformanceDate
|  | EndDate | 2022-02-10 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:endPerformanceDate
|  | Performance location | laboratory | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:performanceLocation
|  | Issue Date | 2022-02-11 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:issueDate
Calibration laboratory | Kalibreringslaboratorium |  |  | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory
|  | CalLab_CompanyName | Dandiag A/S | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:companyName
|  | CalLab_Email | nln@dandiag.com | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:eMail
|  | CalLab_Phone | 2334 45 56 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:phone
|  | CalLab_City | Ishøj | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:location/dcc:city
|  | CalLab_Country | DK | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:location/dcc:countryCode
|  | CalLab_PostalCode | 2635 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:location/dcc:postCode
|  | CalLab_Street1 | Baldershøj | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:location/dcc:street
|  | CalLab_Street_No | 19 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:location/dcc:streetNo
|  | CalLab_Webpage |  | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:location/dcc:further/dcc:content
|  | SignedBy | Erling T. Nielsen | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:respPersons/dcc:respPerson/dcc:person/dcc:name/dcc:content
Customer | Kunde |  |  | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer
|  | Customer_CompanyName | Dandiag A/S | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:companyName
|  | Customer_Att | Scott Kenney Hansen | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:attPerson
|  | Customer_Email | skh@dandiag.com | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:eMail
|  | Customer_Phone | 23 34 3456 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:phone
|  | Customer_City | Ishøj | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:city
|  | Customer_Country | DK | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:countryCode
|  | Customer_PostalCode | 2635 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:postCode
|  | Customer_CompanyStreet1 | Baldershøj | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:street
|  | Customer_CompanyStreet_No | 19 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:streetNo
|  | Customer_Webpage |  |  | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:further/dcc:content

## Statements
Include in DCC | category | id | heading lang1 | body lang1 | heading lang2 | body lang2
x | general | s01 | Reprint | Excerpts of this report may only be reproduced with the written permission of the calibration laboratory. | Gengivelse af rapport i uddrag | Denne rapport må ikke gengives i uddrag undtaget med laboratoriets skriftlige tilladelse.
x | general | None
x | accreditation | acc1 | Accreditation | The calibration is caried out under DANAK accreditation no..490 unless clearly statet in the tabel of results. DANAK is the national accreditation body in Denmark in compliance with Regulation (EC) No. 765/2008 of the European Parliament and of the Council. DANAK is covered by the multilateral agreements for testing, medical examination, calibration, proficiency testing providers and reference material producers under European co-operation for Accreditation (EA) and under International Laboratory Accreditation Cooperation (ILAC) based on peer-evaluation. This implies that accredited reports and certificates issued by companies accredited by DANAK are recognized across borders by members of EA and ILAC equal to reports and certificates issued by companies accredited by these members | Akkreditering | Kalibreringen er udført under DANAK akkreditering nr. 490 med mindre det tydeligt fremgår i resultattabellerne. DANAK er det nationale akkrediteringsorgan i Danmark i overensstemmelse med Europa-Parlamentet og Rådetsforordning (EF) Nr. 765/2008. DANAK er omfattet af de multilaterale aftaler for prøvning, medicinsk undersøgelse, kalibrering, udbud af præstationsprøvning og produktion af referencematerialer i European co-operation for Accreditation (EA) og i International Laboratory Accreditation Cooperation (ILAC) baseret på peer-evaluering. Dette indebærer, at akkrediterede rapporter og certifikater udstedt af virksomheder akkrediteret af DANAK anerkendes på tværs af landegrænser af medlemmer i EA og ILAC på linje med rapporter og certifikater udstedt af disse medlemmers akkrediterede virksomheder.
x | method | meth1 | Calibration method | Calibration is performed in accordance with Dandiags internal procedure PIP-005 | Kalibrering metode | Pipetten er kalibreret med en multikanalsvægt. Dandiags interne procedure PIP-005
x | method | meth2 | Used method | The calibration is carried out by using the gravimetric performance test method. The used method is based on the DS/EN ISO 8655 standard. The used liquid is water according to DS/EN ISO 3696, grade 3. The conversion from mass to volume is done by using the calculation in DS/ISO/TR 20461. | Anvendt metode | Kalibreringen er udført ved brug af den gravimetriske performance-test-metode. Metoden er baseret på standarden DS/EN ISO 8655. Der er anvandt fydende vand af grad3 jvf. DS/EN ISO 3696. Konvertering fra masse til volumen er udført I overensstemmelse med DS/ISO/TR 20461.
x | method | meth3 | Z-factor | Z-factor is calculated according to DS/EN ISO 8655 part 6. | Z-faktor | Z-faktor er beregnet ifølge DS/EN ISO 8655 part 6.
x | method | meth4 | Uncertainty | The reported expanded uncertainty [±U] is given as the measurement standard uncertainty multiplied with the coverage factor k=2 which corresponds to a coverage probability of approximately 95% | Usikkerhed | Den rapporterede måleusikkerhed er angivet som standard måleusikkerheden ganget med dækningsfaktoren k=2 svarende til en dækningssansylighed på ca. 95%.
x | method | env1 | Environment | Pressure: 1011 hPa ±10 hPa | Miljøfaktor | Tryk: 1011 hPa ±10 hPa
x | method | env2 | Environment | Temperatur: 20,5°C±0,4°C | Miljøfaktor | Temperatur: 20,5°C±0,4°C
x | method | env3 | Evironment | relative humidity: 59,4 %RH ±10 %RH | Miljøfaktor | Relativ luftfugtighed: 59,4 %RH ±10 %RH

## Settings
include in dcc | id | heading lang1 | body lang1 | heading lang2 | body lang2 | parameter | value | unit | softwareInstruction
x | c1 | Channel |  | Kanal |  | Channel | 1
x | c2 | Channel |  | Kanal |  | Channel | 2
x | c3 | Channel |  | Kanal |  | Channel | 3
x | sin | Speed in | Same speed is used in all measurements | Hastighed ind | Samme hastighed er anvendt I alle målinger | Speed | 5
x | sout | speed out | Same speed is used in all measurements | Hastighed ud | Samme hastighed er anvendt I alle målinger | Speed | 5

## Accreditation
Include in DCC | id | accreditationCountry | accreditationBody | accreditationLabId | accreditationNorm | accreditationApplicability
x | acc0 | DK | DANAK | 490 | 17025 | 1

## Equipment
Inclued in dcc | category | id | heading lang1 | heading lang2 | manufacturer | productName | productNumber | id1 | id1 issuer | id1 heading lang1 | id1 heading lang2 | id2 | id2 issuer | id2 heading lang1 | id2 heading lang2
x | item | item1 | Pipette with 12 channels 50-1200 µm | Pipette med 12 kanaler 50-1200 µm | Biohit | Picus |  | ID107 | owner | Customer iD | Kundens ID | 12008522 | manufacturer | Serial number | Serienummer
x | item | item2 | Pipette tip (1200 µl) | Pipettespids | Satorius | Optifit |  | - | owner | - | - | ss23987 | manufacturer | Lot number | Lot-nummer
x | reference | ref1 | Balance with 0.00001 g resolution last calibrated on 2021-11-25. | Vægt er med en opløsning på 0,00001 g. Kalibreret d. 2021-11-25 | Satorius |  |  | ID148 | owner | Laboratory's ID for the used reference | Laboratoriets ID nummer | - | manufacturer | - | -

## CalTableC1
DCCTable |  | heading lang1 | heading lang2
tableId | caltab1 | Calibration result chanel 1 | Kalibreringsresultat kanal 1
itemRef | item1 item2
settingRef | c1
numRows | 3
numColumns | 13
scope | - | reference | reference | itemIndication | itemBias | itemBias | itemBias | itemBias | itemBias | itemBias | itemBias | environment | environment
dataCategory | - | TargetValue | Value | Value | Value | ExpandedUncertainty | Value | ExpandedUncertainty | UncertaintyCoverageFactor_k | repeatability | repeatability | Value | Value
measurand | - | Measure.Volume | Measure.Volume | Measure.Volume | Measure.Volume | Measure.Volume | Measure.Volume | Measure.Volume | Measure.Volume | Measure.Volume | Measure.Volume | miscelaneous01 | miscelaneous01
unit | - | \micro\litre | \micro\litre | \micro\litre | \micro\litre | \micro\litre | \percent | \percent | - | \micro\litre | \percent | \one | \one
metaDataType | customerTag | Data | Data | Data | Data | Data | Data | Data | Data | Data | Data | Data | Statement
humanHeading lang1 | Customer Tags | Calibration Point | Reference Value | Item indication | Systematic error | Calibration uncertainty | Systematic error (relative) | Calibration uncertainty (relative) | Uncertainty k factor | Repeatability | Repeatability (relative) | Z-factor | Calculation of Z-factor
humanHeading lang2 | Kundens kode | Kalibreringspunkt | Referencens værdi | Emnets visning | Systematisk fejl | Kalibreringsusikkerhed | Systematisk fejl (relativ) | Kalibreringsusikkerhed (relativ) | Usikkerhed k-faktor | Repeterbarhed | Repeterbarhed (relativ) | Z faktor | Bergning af Z-faktor
data_row1 | p1 | 120 | 121.21 | 120 | 1.21 | 0.35 | 1.01 | 0.29 | 2 | 0.13 | 0.11 | 1.003 | meth3
data_row2 | p2 | 600 | 600.3 | 600 | 0.3 | 1.5 | 0.05 | 0.25 | 2 | 0.24 | 0.04 | 1.003 | meth3
data_row3 | p3 | 1200 | 1199.1 | 1200 | -0.9 | 2.4 | -0.08 | 0.2 | 2 | 0.48 | 0.04 | 1.003 | meth3

## MS120TableC1
DCCTable |  | heading lang1 | heading lang2
tableId | MS120 | Series of measurement channel 1 | Måleserie kanal 1
itemRef | item1 item2
settingRef | c1
numRows | 10
numColumns | 3
scope | - | reference | reference
dataCategory | - | TargetValue | Value
measurand | - | Measure.Volume | Measure.Volume
unit | - | \micro\litre | \micro\litre
metaDataType | customerTag | Data | Data
humanHeading lang1 | Customer Tags | Calibration Point | Reference value
humanHeading lang2 | Kundens kode | Kalibreringspunkt | Referenceværdi
data_row1 | p1 | 120 | 121.3229
data_row2 | p2 | 120 | 121.1223
data_row3 | p3 | 120 | 121.1724
data_row4 | p4 | 120 | 121.1022
data_row5 | p5 | 120 | 121.2025
data_row6 | p6 | 120 | 121.1825
data_row7 | p7 | 120 | 121.3129
data_row8 | p8 | 120 | 121.042
data_row9 | p9 | 120 | 121.1724
data_row10 | p10 | 120 | 121.5034

## MS600TableC1
DCCTable |  | heading lang1 | heading lang2
tableId | MS600 | Series of measurement channel 1 | Måleserie kanal 1
itemRef | item1 item2
settingRef | c1
numRows | 10
numColumns | 4
scope | - | reference | reference | itemIndication
dataCategory | - | TargetValue | Value | Value
measurand | - | Measure.Volume | Measure.Volume | Measure.Volume
unit | - | \micro\litre | \micro\litre | \micro\litre
metaDataType | customerTag | Data | Data | Data
humanHeading lang1 | Customer Tags | Calibration Point | Reference value | Item indication
humanHeading lang2 | Kundens kode | Kalibreringspunkt | Referenceværdi | Emnets visning
data_row1 | p1 | 600 | 600.5463 | 600
data_row2 | p2 | 600 | 600.466 | 600
data_row3 | p3 | 600 | 599.784 | 600
data_row4 | p4 | 600 | 600.5262 | 600
data_row5 | p5 | 600 | 600.4058 | 600
data_row6 | p6 | 600 | 600.1049 | 600
data_row7 | p7 | 600 | 600.3457 | 600
data_row8 | p8 | 600 | 600.3757 | 600
data_row9 | p9 | 600 | 600.7067 | 600
data_row10 | p10 | 600 | 600.1952 | 600

## MS1200TableC1
DCCTable |  | heading lang1 | heading lang2
tableId | MS1200 | Series of measurement channel 1 | Måleserie kanal 1
itemRef | item1 item2
settingRef | c1
numRows | 10
numColumns | 4
scope | - | reference | reference | itemIndication
dataCategory | - | TargetValue | Value | Value
measurand | - | Measure.Volume | Measure.Volume | Measure.Volume
unit | - | \micro\litre | \micro\litre | \micro\litre
metaDataType | customerTag | Data | Data | Data
humanHeading lang1 | Customer Tags | Calibration Point | Reference value | Item indication
humanHeading lang2 | Kundens kode | Kalibreringspunkt | Referenceværdi | Emnets visning
data_row1 | p1 | 1200 | 1199.488 | 1200
data_row2 | p2 | 1200 | 1199.046 | 1200
data_row3 | p3 | 1200 | 1198.043 | 1200
data_row4 | p4 | 1200 | 1198.615 | 1200
data_row5 | p5 | 1200 | 1199.478 | 1200
data_row6 | p6 | 1200 | 1199.949 | 1200
data_row7 | p7 | 1200 | 1198.966 | 1200
data_row8 | p8 | 1200 | 1199.087 | 1200
data_row9 | p9 | 1200 | 1198.896 | 1200
data_row10 | p10 | 1200 | 1199.247 | 1200

## Software (hidden)
swID | name | release | type | description | content | file | formula

## Locations (hidden)
| heading lang1 | heading lang2 | Id | Category | CompanyName | Att | Email | Phone | City | Country | PostalCode | Street1 | Street2 | Street_No | Webpage
heading lang1 |  |  |  |  | Company name |  | email: | Phone: | City: | Country code | Postal Code
heading lang2 |  |  |  |  | Firmanavn |  | email: | Tel.: | By: | Landekode: | Postnr.:
x | Calibration laboratory | Kalibreringslaboratoriet | lab | laboratory | DFM A/S |  | Administration@dfm.dk | +45 7730 5800 | Hørsholm | DK | 2970 | Kogle Alle |  | 5 | www.dfm.dk
x | Client | Kunde | cust | customer | Danish Measurement Company Ltd. | Mads Målermand | mm@dmc.dk | 34507624 | Måløv | DK | 9899 | Målervej |  | 16, 1. sal