## Definitions
scopeTypes | dataCategoryType | measurandType | unit | equipmentClass
dataInfo | Value | Measure.Capacitance | \metre
reference | TargetValue | Measure.Conductance | \kilogram
itemIndication | ToleranceLimitUpper | Measure.Conductivity | \second
itemBias | ToleranceLimitUpper | Measure.Current.AC | \ampere
| AcceptanceLimitUpper | Measure.Current.AC.Sinewave | \kelvin
| AcceptanceLimitUpper | Measure.Current.AC.Squarewave | \mole
| Conformity | Measure.Current.AC.Trianglewave | \candela
| UsedReference | Measure.Current.DC | \one
| UsedMethod | Measure.Density.Mass.Gas | \day
| UsedEquipment | Measure.Density.Mass.Liquid | \hour
| repeatability | Measure.Density.Mass.Solid | \minute
| Other | Measure.Force | \degree
| customerTag | Measure.Frequency | \arcminute
| laboratoryTag | Measure.Frequency.AmplitudeModulation.Rate | \arcsecond
| accreditationApplies | Measure.Frequency.FrequencyModulation.Deviation | \gram
| ExpandedUncertainty | Measure.Frequency.FrequencyModulation.Rate | \radian
| UncertaintyCoverageFactor_k | Measure.Frequency.PhaseModulation.Rate | \steradian
| UncertaintyCoverageProbability | Measure.Humidity.Absolute | \hertz
|  | Measure.Impedance | \newton
|  | Measure.Inductance | \pascal
|  | Measure.Length | \joule
|  | Measure.Length.Circumference | \watt
|  | Measure.Length.Diameter | \coulomb
|  | Measure.Length.Form.Flatness | \volt
|  | Measure.Length.Form.Parallelism | \farad
|  | Measure.Length.Form.Perpendicularity | \ohm
|  | Measure.Length.Form.Roughness | \siemens
|  | Measure.Length.Form.Roundness | \weber
|  | Measure.Length.Form.Sphericity | \tesla
|  | Measure.Length.Form.Straightness.Axis | \henry
|  | Measure.Length.Form.Straightness.Surface | \degreecelsius
|  | Measure.Length.Radius | \lumen
|  | Measure.Mass.Apparent | \lux
|  | Measure.Mass.Conventional | \becquerel
|  | Measure.Mass.True | \sievert
|  | Measure.Phase.PhaseModulation | \gray
|  | Measure.Phase.ReflectionFactor.RF | \katal
|  | Measure.Phase.TransmissionFactor | \hectare
|  | Measure.PhaseNoise.SideBand | \litre
|  | Measure.Power.RF.Sinewave | \tonne
|  | Measure.Pressure.Hydraulic.Static | \electronvolt
|  | Measure.Pressure.Pneumatic.Absolute.Static | \dalton
|  | Measure.Pressure.Pneumatic.Differential.Static | \astronomicalunit
|  | Measure.Pressure.Pneumatic.Gage.Static | \neper
|  | Measure.Ratio.AmplitudeModulation | \bel
|  | Measure.Ratio.Density.Mass | \decibel
|  | Measure.Ratio.Humidity.Relative | \deca
|  | Measure.Ratio.Humidity.Specific | \hecto
|  | Measure.Ratio.Power.ReflectionFactor.RF | \kilo
|  | Measure.Ratio.Power.RF.Sinewave.Delta.Frequency | \mega
|  | Measure.Ratio.Power.RF.Sinewave.Delta.Power | \giga
|  | Measure.Ratio.Power.TransmissionFactor | \tera
|  | Measure.Ratio.Torque | \peta
|  | Measure.Ratio.Voltage.AC.Ripple.OnDC | \exa
|  | Measure.Ratio.Voltage.AC.Sinewave.Delta.Frequency | \zetta
|  | Measure.Ratio.Voltage.AC.Sinewave.Delta.Voltage | \yotta
|  | Measure.Resistance | \deci
|  | Measure.Temperature | \centi
|  | Measure.Temperature.Radiometric | \milli
|  | Measure.Temperature.Simulated.PRT | \micro
|  | Measure.Temperature.Simulated.RTD | \nano
|  | Measure.Temperature.Simulated.Thermocouple | \pico
|  | Measure.Time.Transition | \femto
|  | Measure.Time.UTC | \atto
|  | Measure.Torque | \zepto
|  | Measure.Torque.HydraulicPressure | \yocto
Measure.Voltage.AC
Measure.Voltage.AC.Ripple.OnDC
Measure.Voltage.AC.Sinewave
Measure.Voltage.AC.Squarewave
Measure.Voltage.AC.Trianglewave
Measure.Voltage.DC
Measure.Weight
Source.Capacitance
Source.Conductance
Source.Conductivity
Source.Current.AC.Sinewave
Source.Current.AC.Squarewave
Source.Current.AC.Trianglewave
Source.Current.DC
Source.Density.Mass.Gas
Source.Density.Mass.Liquid
Source.Density.Mass.Solid
Source.Force
Source.Frequency.AC.Squarewave
Source.Frequency.Arbitrary.Cardiograph
Source.Humidity.Absolute
Source.Impedance
Source.Inductance
Source.Length
Source.Length.Circumference
Source.Length.Diameter
Source.Length.Form.Perpendicularity
Source.Length.Form.Roundness
Source.Length.Form.Sphericity
Source.Length.Form.Straightness.Surrface
Source.Length.Radius
Source.Mass.Apparent
Source.Mass.Conventional
Source.Mass.True
Source.Period.Marker
Source.Period.Squarewave
Source.Power.AC.Sinewave
Source.Power.AC.Sinewave.Simulated
Source.Power.DC.Simulated
Source.Power.Noise.Terminated
Source.Power.RF.Sinewave
Source.Pressure.Hydraulic.Static
Source.Pressure.Pneumatic.Absolute.Static
Source.Pressure.Pneumatic.Differential.Static
Source.Pressure.Pneumatic.Gage.Static
Source.Ratio.AmplitudeModulation
Source.Ratio.Humidity.Relative
Source.Ratio.Humidity.Specific
Source.Ratio.Power.RF.Sinewave.Delta.Frequency
Source.Ratio.Voltage.AC.Sinewave.Delta.Frequency
Source.Ratio.Voltage.AC.Sinewave.Delta.Voltage
Source.Resistance
Source.Temperature
Source.Temperature.FixedPoint
Source.Temperature.Radiometric
Source.Temperature.Simulated.PRT
Source.Temperature.Simulated.RTD
Source.Temperature.Simulated.Thermocouple
Source.Time.Marker
Source.Torque
Source.Voltage.AC.Sinewave
Source.Voltage.DC
Source.Voltage.DC.Delta.Voltage
Source.Voltage.Noise.Terminated
Source.Voltage.Shorted
Source.Weight
temperatureAbsolute
temperatureDifference
pressureInternal
pressureAbsolute
pressureExcess
geometryLength
geometryDiameter
geometryFormError
geometryXcoord
geometryYcoord
geometryZcoord
volumeAbsolute
volumeRelative
diameter
roundness
power
resistance
customisedMesurand
metaData

## AdministrativeData
UserLabel | Template Label | Value | XPATH
| DFM_Software_Name | DFM substituter | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:dccSoftware/dcc:software/dcc:name/dcc:content
| DFM_Software_Version | v0.0.1 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:dccSoftware/dcc:software/dcc:release
| DFM_Unique_ID | DFM-T220000 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:uniqueIdentifier
Certificate_no | DFM_CertificateNo | DFM-T220000 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:identifications/dcc:identification[1]/dcc:value
| DFM_OrderNo | 8000650430 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:identifications/dcc:identification[2]/dcc:value
| DFM_ArrivalDate | 2022-02-18 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:receiptDate
| DFM_StartDate | 2022-02-21 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:beginPerformanceDate
Calibr_date | DFM_EndDate | 2022-02-22 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:coreData/dcc:endPerformanceDate
| DFM_Item1_Description | Digitaltermometer | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:items/dcc:name/dcc:content[1]
| DFM_Item1_Name | Fluke | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:items/dcc:item/dcc:name/dcc:content[1]
| DFM_Item1_Maufacturer | Temperature sensor | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:items/dcc:item/dcc:manufacturer/dcc:name/dcc:content
| DFM_Item1_Model | 1523 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:items/dcc:item/dcc:model
| DFM_Item1_SerialNo | 8864333 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:items/dcc:item/dcc:identifications/dcc:identification[1]/dcc:value
Id_no | DFM_Item1_ID | DMC-2769 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:items/dcc:item/dcc:identifications/dcc:identification[2]/dcc:value
Supplier_Name | DFM_CalLab_CompanyName | DFM A/S | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:contact/dcc:name/dcc:content
| DFM_CalLab_Email | Administration@dfm.dk | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:contact/dcc:eMail
| DFM_CalLab_Phone | +45 7730 5800 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:contact/dcc:phone
| DFM_CalLab_City | Hørsholm | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:contact/dcc:location/dcc:city
| DFM_CalLab_Country | DK | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:contact/dcc:location/dcc:countryCode
| DFM_CalLab_PostalCode | 2970 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:contact/dcc:location/dcc:postCode
| DFM_CalLab_Street1 | Kogle Alle | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:contact/dcc:location/dcc:street
| DFM_CalLab_Street_No | 5 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:contact/dcc:location/dcc:streetNo
| DFM_CalLab_Webpage | www.dfm.dk | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:calibrationLaboratory/dcc:contact/dcc:location/dcc:further/dcc:content
| DFM_SignedBy | J. S. Nielsen | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:respPersons/dcc:respPerson/dcc:person/dcc:name/dcc:content
| DFM_Customer_CompanyName | Danish Measurement Company Ltd. | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:name/dcc:content
| DFM_Customer_Email | mm@dmc.dk | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:eMail
| DFM_Customer_City | Måløv | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:city
| DFM_Customer_Country | DK | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:countryCode
| DFM_Customer_PostalCode | 9899 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:postCode
| DFM_Customer_CompanyStreet1 | Målervej | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:street
| DFM_Customer_CompanyStreet_No | 16, 1. sal | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:streetNo
| DFM_Customer_Att | Mads Målermand | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:customer/dcc:location/dcc:further/dcc:content[1]
| DFM_Accreditation_CALRegNo | DK-255 | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:statements/dcc:statement[1]/dcc:reference
| DFM_Accreditation_traceable | true | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:statements/dcc:statement[1]/dcc:traceable
| DFM_Accreditation_valid | true | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:statements/dcc:statement[1]/dcc:valid
| DFM_Accreditation_Statement | The calibration is performed under DANAK accreditation no. 255.\nDANAK is one of the signatories ot the EA Multilateral Agreement and the ILAC multilateral agreement for the mutual recognition of calibration certificates. | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:statements/dcc:statement[1]/dcc:declaration/dcc:content[1]
| DFM_Accreditation_Body | DANAK | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:statements/dcc:statement[1]/dcc:respAuthority/dcc:name/dcc:content[1]
| DFM_TermsOfUse_Statement | This certificate may not be reproduced except in full without the written approval of the laboratory. | /dcc:digitalCalibrationCertificate/dcc:administrativeData/dcc:statements/dcc:statement[2]/dcc:declaration/dcc:content[1]
| DFM_Item1_Result1_Name_en | Temperatures accordign to ITS-90 | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:name/dcc:content[2]
| DFM_Item1_Result1_Name_da | Temperaturer i henhold til ITS-90 | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:name/dcc:content[3]
| DFM_Procedure | DFM proceedure Q2KAL251 | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:usedMethods/dcc:usedMethod[1]/dcc:description/dcc:content
| DFM_Reference1_name | Standard Platinum Resistance Thermometer | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:measuringEquipments/dcc:measuringEquipment/dcc:name/dcc:content[2]
| DFM_Reference1_manufacturer | Rosemount | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:measuringEquipments/dcc:measuringEquipment/dcc:manufacturer/dcc:name/dcc:content
| DFM_Reference1_model | 162CE | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:measuringEquipments/dcc:measuringEquipment/dcc:model
| DFM_Reference1_serial | 5091 | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:measuringEquipments/dcc:measuringEquipment/dcc:identifications/dcc:identification[1]/dcc:value
| DFM_Reference1_ID | DFM-1546 | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:measuringEquipments/dcc:measuringEquipment/dcc:identifications/dcc:identification[2]/dcc:value
| DFM_Reference_Statement | Temperature is measured via DFM resistancethermometer Rosemount CE162/5091, with DFM calibration certificate number TIR2015. The calibration is executed in situ. | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:measuringEquipments/dcc:measuringEquipment/dcc:description/dcc:content
| DFM_LabTemperature_Description | The calibrations are made at constant temperatures by comparison wiht laboratory standards traceable to SI. | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:influenceConditions/dcc:influenceCondition/dcc:description/dcc:content[2]
| DFM_LabTemperature_Value | 23 | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:influenceConditions/dcc:influenceCondition/dcc:data/dcc:quantity/si:hybrid/si:real/si:value
| DFM_LabTemperature_Unit | \degreeCelcius | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:influenceConditions/dcc:influenceCondition/dcc:data/dcc:quantity/si:hybrid/si:real/si:unit
| DFM_LabTemperature_U | 2 | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:influenceConditions/dcc:influenceCondition/dcc:data/dcc:quantity/si:hybrid/si:real/si:expandedUnc/si:uncertainty
| DFM_Item1_Norm | DKD-R 5-1:2018 | /dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:usedMethods/dcc:usedMethod[1]/dcc:norm
/dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:results/dcc:result/dcc:data/dcc:list/dcc:quantity[1]/dcc:measurementMetaData/dcc:metaData/dcc:data/dcc:quantity/si:hybrid/si:realListXMLList[2]/si:valueXMLList
/dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:results/dcc:result/dcc:data/dcc:list/dcc:quantity[2]/si:hybrid/si:realListXMLList[1]/si:expandedUncXMLList/si:uncertaintyXMLList
/dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:results/dcc:result/dcc:data/dcc:list/dcc:quantity[2]/si:hybrid/si:realListXMLList[1]/si:expandedUncXMLList/si:coverageFactorXMLList
/dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:results/dcc:result/dcc:data/dcc:list/dcc:quantity[2]/si:hybrid/si:realListXMLList[2]/si:valueXMLList
/dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:results/dcc:result/dcc:data/dcc:list/dcc:quantity[2]/si:hybrid/si:realListXMLList[2]/si:expandedUncXMLList/si:uncertaintyXMLList
/dcc:digitalCalibrationCertificate/dcc:measurementResults/dcc:measurementResult/dcc:results/dcc:result/dcc:data/dcc:list/dcc:quantity[2]/si:hybrid/si:realListXMLList[2]/si:expandedUncXMLList/si:coverageFactorXMLList

## Software
swID | name | release | type | description | content | file | formula

## Statements
id | refID | name | description | countryCodeISO3166_1 | convention | traceable | norm | reference | declaration | valid | date | preiod | respAuthority | data | nonSIDefinition | nonSIUnits | location

## Items
id | Customer ID | equipmentClass | description | swRef | Manufacturer | Product Name | Product number | Serial number

## Table2
DCCTable
TableID | TempCal
refID | itemID1
numRows | 5
numColumns | 12
scope | dataInfo | reference | reference | reference | itemIndication | itemBias | itemBias | itemBias | itemBias | itemBias | itemIndication | dataInfo
dataCategory | customerTag | TargetValue | Value | Value | Value | Value | ExpandedUncertainty | UncertaintyCoverageFactor_k | AcceptanceLimitUpper | Conformity | Value | accreditationApplies
measurand | metaData | temperatureAbsolute | temperatureAbsolute | temperatureAbsolute | temperatureAbsolute | temperatureDifference | temperatureAbsolute | temperatureAbsolute | temperatureAbsolute | temperatureDifference | resistance | metaData
unit | nan | \degreecelcius | \kelvin | \degreecelcius | \degreecelcius | \degreecelcius | \degreecelcius | \degreecelcius | \degreecelcius | \degreecelsius | \ohm | nan
humanHeading | Customer Tags | Calibration Point | Reference Temperature | Reference Temperature in instrument units | Instrument indication | Instrument Error | Uncertainty | k | Upper acceptance limit for item bias | Conformaty statement | Indicated resistance of the calibration item | Accreditation applies
data_row1 | T0_0_degC | 0 | 273.16 | 0.01 | -0.1 | -0.1 | 0.05 | 2 | 0.2 | pass | 1 | y
data_row2 | T1_15_degC | 15 | 288.16 | 15.01 | 15.1 | 0.1 | 0.05 | 2 | 0.2 | pass | 2 | y
data_row3 | T2_25_degC | 25 | 298.17 | 25.02 | 25.2 | 0.2 | 0.05 | 2 | 0.2 | pass | 3 | n
data_row4 | T3_15_degC | 15 | 288.16 | 15.01 | 15.3 | 0.3 | 0.05 | 2 | 0.2 | fail | 4 | y
data_row5 | T4_0_degC | 0 | 273.16 | 0.01 | 0 | 0 | 0.05 | 2 | 0.2 | pass | 5 | y